## Slide 1
Next Generation HbA1c Analysis Using Capillary Electrophoresis

## Slide 2
world leader
dedicated to clinical diagnostic electrophoresis
since 50 years

## Slide 3
Based in South of Paris, France (L. De Vinci Tech Park)
600 employees worldwide
Total surface 43 000 m2
5 Units over 15 000 m2
R&D and production for Chemistry, softwares, hardware…

## Slide 4: Key dates
SEBIA is established
Manufacturing unit established
Creation of the first SEBIA subsidiaries
  SEBIA GmbH
  SEBIA Benelux
  SEBIA Inc
  SEBIA Italia
  SEBIA Hispania
SEBIA moves to the EVRY BIOPARC
New subsidiaries:
  SEBIA UK
  SEBIA China
  SEBIA Brasil
Acquisition of Beckman electrophoresis activity
Introduction of cellulose acetate
CELLOMATIC
CELLOSYSTEM
Launch of the HYDRAGEL range
HYDRASYS
CAPILLARYS
Great expansion of the capillary electrophoresis .
The range of analysis possibilities grows steadily
MINICAP
GelSCAN Technology
CAPILLARYS Flex Piercing/ASSIST
Hb Whole Blood…
Minicap Flex
1967
1982
1986
2001
1994
2004
2006
2007
2008
2010
2011
Launch of Capillarys HbA1c

## Slide 8: Capillary Electrophoresis Principle

## Slide 10: What goes on inside the capillary?
Anode
--
--
--
---
---
---
-
-
-
Cathode
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Detection
Injection
Electro-osmotic flow
Electrical field

## Slide 11
Detector:
Simultaneous measurements (CAPILLARYS - 8; MINICAP - 2)
Capillary
Optical fibers
Peltier device
Temperature control

## Slide 12: Basic CE Diagram
Notes:   Electroosmotic Flow moves the buffer and sample forward because the charge in the buffer is negative.
  During the flow, higher power + molecules flow to the front of the pack.
  As the molecules move past the detector, they are now grouped as to their relative charge and type.
  This methodolgy for performing electrophoresis is consider the most accurate because there is no interference with the medium such as in paper or gel.
  There is no stain affinity issue.

## Slide 13
HEMOGLOBIN INVESTIGATION
by
CAPILLARY
ELECTROPHORESIS
HbA
HbF
HbS
HbA2
HbC

## Slide 14: Separation of normal hemoglobins by capillary electrophoresis
β1
β2
α1
α2
α1
α2
α1
α2
γ1
γ2
δ1
δ2
Hb A > 96,5%
Hb F < 1%
Hb A2 < 3,5%

## Slide 15
Hb A, F and A2 are detected automatically
Hemoglobin electrophoresis on CE systems

## Slide 16: Hb S glycated and Hb A2
HPLC
SEBIA CAPILLARYS 2
Notes: It is well-known and published that with HPLC technology, in the presence of Hb S, Hb A2 is falsely elevated due to Hb S adducts such as carbamylated Hb S and/or glycated Hb S. Pathologists must take caution in interpreting such Hb A2 results and not assume that the patient is also beta-thalassemic.
Using Capillary electrophoresis, all glycated fractions co-migrate with the main corresponding peak and are not separated. In the presence of Hb S, CE technology will provide the user with a more accurate Hb A2 quantitation and is a better indicator of potential beta-thalassemia.

## Slide 17: Hb E and Hb A2
HPLC
SEBIA CAPILLARYS 2
Notes: Hb E coelutes with Hb A2 on HPLC. Either way, Hb E and Hb A2 cannot be quantified separately with the use of HPLC technology.

## Slide 18: Hb H & Hb Bart’s
HPLC
SEBIA CAPILLARYS 2
Notes: With CE, alpha-thalassemia products (Hb H & Hb Bart’s) are clearly separated and detected. Using HPLC, Hb Bart’s is hardly seen and notated by the user, and, due to the fast elution of Hb H (Hb H and Bart’s might elute “prior to the start of integration with HPLC”), it is either not seen or coelutes with Hb Bart’s and therefore is not notated here by the user. Also, with HPLC, Hb Constant Spring can easily be missed.

## Slide 19: What is HbA1c?

## Slide 20: HbA1c: What are we looking for?
β1
β2
α1
α2
α1
α2
α1
α2
γ1
γ2
δ1
δ2
Hb A > 96,5%
Hb F < 1%
Hb A2 < 3,5%

## Slide 21: HbA1c: What are we looking for?
Hb A0
93-95%
Hb A1 = GHb
Glycated Hbs
5-7%
Hb A
+
+
+
Hb A1a
0,5%
Fructose-1,6-diphosphate
Glucose-6-phosphate
Hb A1b
0,5%
pyruvate
Hb A1c
4-6%
glucose
Hb F
Hb A2
Glycation at the
N-terminal Valin
of the β-globin chain

## Slide 22: IFCC : Reference method
In 2002, the International Federation of Clinical Chemistry (IFCC) designs a new reference method for specific measurement of concentration of HbA1c
IFCC (mmol/mol) and NGSP (%) values are correlated through a specific formula
All commercially available techniques are calibrated based on the material produced by IFCC

## Slide 23
1. Digest, Glu-C
2. Separate, LC
3. Detect, MS
Zorbax SB-CN,
2.1 x 150 mm
50 °C, 300 l/min
[M+2H]+2
1 mg Hb
pH 4.3, 37 °C, 18 hours
IFCC LC/MS Reference Method
IFCC Reference METHOD #2: LC/CE
5. Detect,
UV 214 nm
3. Collect, concentrate
4. Separate, CE

## Slide 24
Typical chromatograms from LC/MS analysis
RT:
7.93 - 15.08
8
9
10
11
12
13
14
15
Time (min)
0
20
40
60
80
100
0
20
40
60
80
100
Relative Abundance
10.83
11.84
NL:
6.00E7
NL:
7.75E6
[M+2H]2+
Native
Hexapeptide
m/z 348.2-348.4
[M+2H]2+
Glycated
Hexapeptide
m/z 429.2-429.4
LC/MS

## Slide 25
Calculation Formula
by the IFCC Reference Methods
%HbA1c
=
HbA1c
(HbA1c + HbA0)

## Slide 26
ION-EXCHANGE HPLC

## Slide 27
Hb A0
Hb A
Hb A1a
Hb A1b
Hb A1c
Hb F
Hb A2
%HbA1c =
HbA1c area / ∑ HbA area
Ion-Exchange: Principle

## Slide 28
are minor glycated HbA
What is really being measured?
the following calculation is applied:
[100 x (HbA1c area / ∑ HbA area)]
What does ∑ HbA area correspond to?
∑ HbA area = ∑ ( HbA1a1/a2 + HbA1b + LA1c + HbA1c + P3 + HbA0 )
What is P3?
P3
P3 = HbA0 Gluthatione & HbA0 oxydated
This fraction is limited to 10% by the software
What is LA1c?
LA1c
LA1c = Labile HbA1c including Carbamylated Hb
What about HbA0?
HbA0 co-elutes with another Hb
= HbA2
What are A1a & A1b?
A1a1 +A1a 2
& A1b
HbA1a1/a2
HbA1b
What is P4?
%HbA1c =
P4 = degradation products?
HbA1c
+ Labile Hb A1c
+ HbA0
+ HbA2 )
+ Hb A1b
[100 x (HbA1c /
so basically
+ HbA1c
+ HbA0 oxydated
+ Hb A1a 2
+ Carbamylated Hb
+ HbA0 Gluthatione
( Hb A1a1
+ Hb F
in which the co-elutes
HbF
HPLC Calculation
[100 x (HbA1c area / (HbA1a1 + HbA1a2 + HbA1b + HbF + Labile A1c + Carbamylated Hb + HbA1c + HbA0 Gluthatione + HbA0 oxydated + HbA0 + HbA2) area) x slope + Intercept]

## Slide 29
NORMAL SAMPLE
/ ∑ HbA area
%HbA1c =
S1A
S1B
LS1C+
S1C
S0
HOMOZYGOUS HbS
HETEROZYGOUS A/S
HbA1c area
-HbS (heterozygous) – HbS appears as an H-V1 peak following the A0 peak and there is no carryover observed in the chromatograms that follow. HbS (heterozygous) does not interfere with quantitation of HbA1c. Report the HbA1c result with the following coded comment: (Abnormal hemoglobin variant observed).
HPLC & Hb Variant

## Slide 30: HbA1c Sebia (Clear Cut and Precise)

## Slide 31
A1a
Hb F
LA1c
Hb A1c
Hb A0
Unknown
Unknown
P3
P4
Hb E, Hb D
?
?
?
?
?
?
?
HPLC
Long elution time
?
Clear-cut separation
A1b
Just keep the essential

## Slide 32
What do we measure?
We apply the following calculation:
%HbA1c =
100 x [HbA1c area / (HbA1c area + HbA0 area)]
What should we measure?
Definition of the Analyte according to IFCC:
We apply the same calculation formula as that of the IFCC Reference Method
Accurate
HbA1c measurement
HbA1c is biochemically characterised as the stable adduct of glucose to the N-terminal amino group of the β-chain of hemoglobin A0
HbA1c =
HbA1c / (HbA1c + HbA0)

## Slide 33
- A reduction of 1% of the HbA1c concentration decrease risk of :
- Infarctus : by 14 %
- Cardiovascular events : 37 %
- Amputations : 43 %
The American Diabetes Association (ADA) recommends cutoff values of:
  Hb A1c = 6,5% (diagnostic)
  HbA1c = 5,7% - 6,4% (prediabetic)
Method must be TRUE AND PRECISE
Clinician must be informed of any hematological conditions before HbA1c result interpretation
Importance of accuracy in HbA1c measurement
33

## Slide 34
Poor trueness
Poor precision
Poor trueness
Good precision
Good trueness
Poor precision
Good trueness
Good precision
Not accurate
Not accurate
Accurate
Not accurate

## Slide 35: Trueness
First evaluation of CAPILLARYS 2 Flex Piercing as a new analyzer
for HbA1c assay by capillary electrophoresis
S. Jaisson, N. Leroy, J. Meurice, E. Guillard, P. Gillery
Clin Chem Lab Med 50(10): 1769-75 (2012)

## Slide 36: Precision
Precision was tested using the CLSI protocol EP5 [ 5 ] and measured for 20 working days, 2 runs a day, with a duplicate per run, and with 10 other samples included in the run) at low (38 mmol/mol) and high (68 mmol/mol) HbA 1c levels. The instrument is equipped with 8 parallel capillaries, and therefore, EP5 was executed for each of the capillaries
HbA1c: Performance of the SEBIA CAPILLARYS 2 Flex Piercing
C. Weykamp, H. Waenink-Wieggers, E. Kemna, C. Siebelder
Clin Chem Lab Med (2012); DOI: 10.1515/cclm-2012-0560

## Slide 38: Factors to consider when looking at an HbA1c result:
Analytical interference:
HbA1c value obtained is not the correct value because the analytical method failed to accurately measure HbA1c
INCORRECT MEASUREMENT
Biological interference:
HbA1c value obtained is the correct value but its interpretation can be misleading because of physiopathological condition
MISINTERPRETATION

## Slide 39: Limitations and Interferences of Hb A1C
Analytical
INCORRECT MEASUREMENT OF THE ANALYTE
Hb concentration
Hemoglobinopathies
Hb F
Carbamylated Hb
Labile Hb
Acetylated Hb
Extra-Analytical
% HbA1C CORRECT
but
CLINICIAN NEEDS TO ASSESS USE OF HbA1C VALUE FOR DIAGNOSIS
Patient’s Glycation Rate of Hb A
Circulating Life Span of the Red Cell

## Slide 40
Sebia Capillarys Flex Piercing Hb A1c method interferences registered by NGSP (no influence of Hb concentration, labile, carbamylated or acetylated Hb)

## Slide 41
Profil atypique
Possible interférence quantitative si Hb F ou variant >15%
HbA1c value is reported with following comments :
«RBC lifespan may be affected in certain patients».
No Hb F concentration limit
HbF >15%
HbA1c and other fractions

## Slide 43
HPLC & Hb variants
HbS
HbC
HbD
HbE
Ann Biol Clin 2009 ; 67 (1) : 55-65

## Slide 44
Capillary Electrophoresis & Hb A/S
1’
2’
3’
4’
5’
6’
7’
8’
9’
HPLC Long elution (3’)
CE (9’ migration time)
HPLC Short elution (1.5’)

## Slide 45
Case 4
Heterozygous A/S
Capi 2FP
HbA1c 5.7%
HPLC (Biorad VII Turbo)
HPLC (Tosoh G7)
HbA1c 3.7%
HbA1c 6.7%

## Slide 46
Capillary Electrophoresis & Hb A/E
HPLC Long elution (3’)
HPLC Short elution (1.5’)

## Slide 47
Case 2
Heterozygous A/E
HPLC (Tosoh G7)
HbA1c 4.4%
Cap2FP
HbA1c 5.9%
HPLC (Variant II Turbo)

## Slide 48
Capillary Electrophoresis & Hb A/D
HPLC Long elution (3’)
HPLC Short elution (1.5’)

## Slide 49
“The CAPILLARYS 2 Flex Piercing did not show any clinically significant interference with any of the S, C, D, E variants evaluated.”
‘’ARKRAY and TOSOH show many clinically significant interferences with variants’’
“The presence of HbS trait produced statistically significant differences for all 4 methods except CAPILLARYS 2 Flex Piercing.”
Conclusions

## Slide 50: Limitations and Interferences of Hb A1C
Analytical
INCORRECT MEASUREMENT OF THE ANALYTE
Hb concentration
Hemoglobinopathies
Hb F
Carbamylated Hb
Labile Hb
Acetylated Hb
Extra-Analytical
% HbA1C CORRECT
but
CLINICIAN NEEDS TO ASSESS USE OF HbA1C VALUE FOR DIAGNOSIS
Patient’s Glycation Rate of Hb A
Circulating Life Span of the Red Cell

## Slide 51: HbA1c reflects the glycation of Hemoglobin A. Since red blood cells in a normal patient have an average lifespan of 120 days, HbA1c reflects mean glycemia for the previous 3 months
Any physiopathological condition altering red blood cell lifespan will lead to a clinical interference

## Slide 52
The Pros and the Cons of using HbA1c for Diabetes Diagnosis
David B.Sacks; AACC Webinar April 10th 2012

## Slide 53: Red Cell Life Span
Hb Variant | Avg. RBC Lifespan in Days (n)
HbSS | 17 (10)
HbSC | 28 (14)
HbCC | 29 (7)
HbS-β Thal | 75 (2)
HbAC | 82 (9)
HbAS | 93 (3)
Wild-type | 120
Lancet. 1943 Blood. 1969 Blood. 1970 Blood. 1975

## Slide 54
HbA1c &
β-Thalassemia

## Slide 55
HA-8180
VARIANT II
G8
HbA2 ?

## Slide 57
Red cell life span in heterozygotes for Beta-thalassemia ranges
from 81 to 93 days

## Slide 58
Beta-thalassemia screening
Alarm (!) if %HbA2 is above 3%

## Slide 59: Iron Deficiency Anemia

## Slide 60
Orange pattern: Elevated HbA1c
Pink pattern: Atypical Profile
Blue pattern: Normal Profile
Mosaic Window: interpretation of 48 results at a glance

## Slide 61: Worklist / Result Table

## Slide 62: Sample Suspended

## Slide 63: Sample Validated

## Slide 64
FLEXIBILITY
LAB 1
LAB 2
OFFICE

## Slide 65
Immunoassays
Boronate Affinity
Ion-exchange HPLC
Capillary Electrophoresis
Evolution of the HbA1c Methods
Robustness
Separation

## Slide 66
High Resolution
Clear-Cut and Precise Separation
Simple IFCC Calculation Formula
Excellent Precision and Accuracy
No Interferences
Thalassemic diagnosis included
Quick View for Release of Results
Limited Maintenance and calibration
40 Samples Per Hour
Auto release of results to LIS
HbA1c

## Slide 67
Thank you